## Slide 1: NHAO NIC Flat
2020-07-28 Yoonsoo P. Bach
All the codes and FITS files (except for the very raw data) are available via my personal Dropbox:
https://www.dropbox.com/sh/s3g5e4dv3m84flx/AACLHo2sRoScz6Fxg3lvbId4a?dl=0

## Slide 2: Combining Dome Flats
Scale each raw frame by its mean value
  multiplied scale is scale[i] = median(image[i])/median(image[0])
CCDCLIP algorithm with 3-sigma clip, median center, nkeep=5.
  Gain, readout noise form header
  The sigma-map is calculated by stddev(survived-values) at each pixel position.
  SNR = combined pixel value / sigma-map
  Normalized flat = combined image / mean(combined image)

## Slide 3: Simple results
Simple combining of all flats
Dust-donuts that were invisible in the flat are visible in the SNR map.
Normalized flat
SNR

## Slide 4: Components of Flats
Header keywords: FILTER, POL-AGL1, INSROT
  Plus: o/e-ray
Notation:
  <FILTER>_<POL-AGL1:04.1f>_<INSROT:+04.0f>_<oe:s>.fits
  j_00.0_+045_e.fits
  = J-filter, HWP 0.0˚, INSROT 45˚, e-ray part.

## Slide 5
Effects of INSROT/POLAGL
The ratio below for many INSROT values will show the effect of INSROT:
Similarly, the ratio below will show the effect of HWP angle (POL-AGL1):

## Slide 6
1. Effect of INSROT
Cut profile at red lines ± 10 pix width are shown.
This is where we usually put our target.
Black dotted lines = min/max at that ± 10 pixel region.
<oe>: <POL-AGL1>
ticks for cut-profiles: 1 ± 0.02 (2%)
ticks for image axis: 100 pix

## Slide 7
1. Effect of INSROT
Cut profile at red lines ± 10 pix width are shown.
This is where we usually put our target.
Black dotted lines = min/max at that ± 10 pixel region.
ticks for cut-profiles: 1 ± 0.02 (2%)
ticks for image axis: 100 pix
<oe>: <POL-AGL1>

## Slide 8
1. Effect of INSROT
Cut profile at red lines ± 10 pix width are shown.
This is where we usually put our target.
Black dotted lines = min/max at that ± 10 pixel region.
ticks for cut-profiles: 1 ± 0.02 (2%)
ticks for image axis: 100 pix
<oe>: <POL-AGL1>

## Slide 9: Notes
The peak-trough locations of dust donuts rotate as INSROT angle.
The dust indicated by orange arrow is visible only in J, while others appear in all o/e rays in all filters.
  Maybe a hint for where the dusts are located.

## Slide 10: 2. Effect of POL-AGL1 (HWP angle)
ticks for cut-profiles: 1 ± 0.01 (1%)
ticks for image axis: 100 pix
Cut profile at red lines ± 10 pix width are shown.
This is where we usually put our target.
Black dotted lines = min/max at that ± 10 pixel region.

## Slide 11: 2. Effect of POL-AGL1 (HWP angle)
ticks for cut-profiles: 1 ± 0.01 (1%)
ticks for image axis: 100 pix
Cut profile at red lines ± 10 pix width are shown.
This is where we usually put our target.
Black dotted lines = min/max at that ± 10 pixel region.

## Slide 12: 2. Effect of POL-AGL1 (HWP angle)
ticks for cut-profiles: 1 ± 0.01 (1%)
ticks for image axis: 100 pix
Cut profile at red lines ± 10 pix width are shown.
This is where we usually put our target.
Black dotted lines = min/max at that ± 10 pixel region.

## Slide 13: Notes
Why there's no dust donut?
  The dust-donuts' shapes are dependent on INSROT, but not on POL-AGL1.
  Hence, a flat with fixed POL-AGL1 (combined over all INSROT) will have no dust donut feature.
  The dust donut, however, will appear in the SNR map.
    Dust = lower ADU = lower SNR (simply considering photon noise)
    As INSROT changes, dust donut's brightness distribution changes  higher stddev  lower SNR
    (SNR := combined value at a pixel / stddev of that pixel)